Document Title: Firmware Design Document (FDD)

Document ID: HS-DES-003 Version: 1.0 Status: DRAFT

Author: Sandhya Patel

Date: September 8, 2025

1.0 Introduction

This document provides the detailed design specifications for the HydroSync Wearable Firmware Subsystem. The firmware is embedded on the ESP32 microcontroller. It describes the software architecture, key functional modules, and data structures necessary to meet the requirements outlined in the Product Requirements Document (HS-REQ-001), adhering to the implementation details derived from the actual codebase.

2.0 Firmware Architecture

The firmware is implemented as a loop-based application within the Arduino IDE environment using the ESP32 BLE library.

2.1 Operational States

State 1: INITIALIZE (Setup)

- On power-up, the firmware initializes serial communication (115200 baud) for debugging.

- It configures I2C communication (SDA: Pin 21, SCL: Pin 22) for the Pulse Oximeter.

- It initializes the Analog-to-Digital Converter (ADC) resolution to 12-bit.

- It attempts to initialize the MAX30100 sensor. If successful, it sets the IR LED current; otherwise, it logs a failure message.

- It initializes the BLE server, creates a Service and Characteristic, and starts advertising the device as "HydroSync_Wearable".

State 2: AWAIT_CONNECTION / IDLE

- The device continuously advertises its presence via BLE until a client (mobile app) connects.

- If a client disconnects, the firmware automatically restarts advertising to allow reconnection.

- A 500ms delay is implemented in the main loop when no device is connected to reduce processing load.

State 3: ACQUIRE_DATA

- The pox.update() function is called frequently to ensure the pulse oximeter executes its low-level signal processing.

- Every 2000ms (2 seconds), the firmware reads:

  - Skin Temperature: Reads analog voltage from Pin 32 (LM35) and converts it to Celsius.

  - GSR (Galvanic Skin Response): Reads raw analog value from Pin 34.

  - Heart Rate: Retrieves the computed heart rate from the MAX30100 driver.

State 4: TRANSMIT_DATA

- If a BLE client is connected, the acquired data is formatted into a single comma-separated string (CSV format).

- The packet is written to the BLE Characteristic and a notification is sent to the client.

- Data is simultaneously printed to the Serial Monitor for debugging.

3.0 Hardware & Pin Configuration

The firmware is configured for the following hardware connections:

| Component | ESP32 Pin | Protocol/Type |
| --- | --- | --- |
| MAX30100 (HR/SpO2) | SDA: 21, SCL: 22 | I2C |
| LM35 (Temperature) | GPIO 32 | Analog (ADC) |
| GSR Sensor | GPIO 34 | Analog (ADC) |

4.0 Key Functional Modules

4.1 Sensor Acquisition Module

- readSkinTemperature(): Reads ADC value from Pin 32, applies a scaling factor (3300.0 / 4095.0 / 10.0), and returns degrees Celsius.

- readGSR(): Reads the raw 12-bit ADC value (0-4095) from Pin 34 representing skin conductance.

- pox.update() & pox.getHeartRate(): Manages the MAX30100 sensor timing and retrieves the calculated beats per minute (BPM).

4.2 BLE Communication Module

- Service UUID: 0000ffe0-0000-1000-8000-00805f9b34fb

- Characteristic UUID: 0000ffe1-0000-1000-8000-00805f9b34fb

- Properties: The characteristic supports NOTIFY, allowing the wearable to push data updates to the phone without the phone requesting them.

- Callbacks: Implements onConnect and onDisconnect to manage the internal deviceConnected flag and restart advertising upon disconnection.

5.0 Data Packet Structure

Unlike the binary structure originally proposed, the actual implementation uses a String-based CSV format for simplicity and ease of parsing on the mobile side.

Packet Format: String

Structure: HeartRate,Temperature,GSR

| Field Index | Data Type (in String) | Example Value | Description |
| --- | --- | --- | --- |
| 0 | Float (1 decimal) | "72.5" | Heart Rate in BPM |
| 1 | Float (2 decimals) | "36.65" | Skin Temperature in Celsius |
| 2 | Integer | "2048" | Raw GSR ADC Value |

Example Output: "72.5,36.65,2048"

6.0 Implementation Notes

- Reporting Frequency: Data is transmitted every 2 seconds (millis() - lastReportTime > 2000).

- ADC Resolution: The code explicitly sets analogReadResolution(12), meaning analog values will range from 0 to 4095.

- Libraries Used:

  - BLEDevice.h, BLEUtils.h, BLEServer.h (ESP32 BLE)

  - Wire.h (I2C communication)

  - MAX30100_PulseOximeter.h (Heart rate processing)