Document Title: Master Verification Plan Document ID: HS-VVP-001 Version: 1.0 Status: DRAFT Author: Sandhya Patel Date: September 8, 2025

1.0 Introduction

This document describes the test procedures to verify that the HydroSync system meets all requirements specified in the Product Requirements Document (HS-REQ-001). Each requirement will be tested, and the results will be recorded. The successful completion of this plan will verify that the system has been built according to its design specifications.

2.0 Verification Test Cases

| Requirement ID | Requirement Description | Test Method / Procedure | Pass/Fail Criteria |
| --- | --- | --- | --- |
| Section 2.1: Hardware Requirements |  |  |  |
| REQ-001 | The wearable device shall incorporate a Galvanic Skin Response (GSR) sensor. | 1. Visually inspect the device's hardware assembly. 2. Run a diagnostic firmware routine to check for a response from the GSR sensor. | The GSR sensor component shall be present and the diagnostic shall report a successful connection. |
| REQ-002 | The wearable device shall incorporate a pulse sensor capable of deriving Heart Rate Variability (HRV). | 1. Visually inspect the device's hardware assembly. 2. Run a diagnostic firmware routine to check for a response from the pulse sensor. | The pulse sensor component shall be present and the diagnostic shall report a successful connection. |
| REQ-003 | The wearable device shall incorporate a temperature sensor to measure peripheral skin temperature. | 1. Visually inspect the device's hardware assembly. 2. Run a diagnostic firmware routine to check for a response from the temperature sensor. | The temperature sensor component shall be present and the diagnostic shall report a successful connection. |
| REQ-004 | The wearable device shall use an ESP32 microcontroller for all data processing and control functions. | 1. Visually inspect the device's main circuit board. | The microcontroller shall be an ESP32 model. |
| REQ-005 | The wearable device shall be powered by a rechargeable Lithium-Ion battery. | 1. Fully charge the device. 2. Run the device in its standard operating mode (streaming data over BLE). 3. Measure the total time until the battery is depleted. | The measured operating time shall be greater than or equal to 5 hours. |
| Section 2.2: Firmware Requirements |  |  |  |
| REQ-006 | The firmware shall acquire data from all integrated sensors at a user-configurable sync interval. | 1. Set the sync interval to 5s in the mobile app. 2. Using a data logger, monitor the BLE data packets for 1 minute. 3. Repeat for a 10s interval. | The data logger shall show packets arriving approximately every 5 seconds (+/- 10%) and every 10 seconds (+/- 10%) respectively. |
| REQ-007 | The firmware shall transmit the acquired sensor data to the paired mobile application using Bluetooth Low Energy (BLE). | 1. Connect the wearable to the mobile app. 2. Observe the live data stream on the app's diagnostic screen for 5 minutes. | Data packets shall be received continuously with no disconnections during the 5-minute test period. |
| Section 2.3: Mobile Application Requirements |  |  |  |
| 2.3.1 Dashboard |  |  |  |
| REQ-008 | The mobile application shall be capable of establishing and maintaining a BLE connection. | 1. Power on the wearable. 2. Open the app and initiate a connection. 3. Leave the app connected for 10 minutes. | The app shall connect to the device within 15 seconds and shall remain connected for the 10-minute duration. |
| REQ-009 | The mobile application dashboard shall display the user's key physiological metrics. | 1. Connect the device. 2. Observe the dashboard. | The dashboard shall display numerical or graphical indicators for Hydration, Fatigue, and Skin Temperature. |
| REQ-010 | The mobile application dashboard shall display a primary status banner. | 1. Connect the device. 2. Observe the dashboard. | A primary status banner (e.g., displaying text like "Optimal" or "Fatigued") shall be present on the dashboard. |
| REQ-011 | The color of the primary status banner shall change to indicate the user's state. | 1. Using a test firmware, send data indicating a "Fatigued" state. 2. Send data indicating a "Healthy" state. | The banner color shall change to red (or equivalent alert color) in the "Fatigued" state and green (or equivalent) in the "Healthy" state. |
| REQ-012 | The mobile application dashboard shall provide navigation buttons. | 1. Observe the dashboard. | Navigation buttons, icons, or tabs for "Trends," "Alerts," and "User Profile" shall be present and visible. |
| 2.3.2 Trends Screen |  |  |  |
| REQ-013 | The Trends screen shall display a graphical plot of the user's historical physiological data over time. | 1. Collect data for 1 hour. 2. Navigate to the Trends screen. | A line graph showing the data from the past hour shall be displayed. |
| REQ-014 | The Trends screen shall provide filter options to view data for individual metrics. | 1. Navigate to the Trends screen. 2. Tap the "Hydration" filter, then the "Fatigue" filter. | Tapping the filters shall correctly show or hide the corresponding data plots on the graph. |
| 2.3.3 Alerts Screen |  |  |  |
| REQ-015 | The Alerts screen shall display a list of all historical alert events. | 1. Using a test firmware, trigger three distinct alerts. 2. Navigate to the Alerts screen. | All three test alerts shall be listed on the screen. |
| REQ-016 | Each alert in the list shall include a description, the measured value, and a timestamp. | 1. Trigger a test "High Temperature" alert with a value of 38.5°C. 2. Navigate to the Alerts screen and inspect the entry. | The alert entry shall contain the text "High Temperature" (or similar), the value "38.5°C", and the correct timestamp. |
| REQ-017 | The Alerts screen shall provide a function for the user to dismiss or clear alerts. | 1. Trigger several test alerts. 2. Navigate to the Alerts screen and tap the "Clear All" button. | All alerts shall be removed from the list. |
| REQ-018 | The Alerts screen shall provide a function to export the alert history as a CSV file. | 1. Trigger several test alerts. 2. Tap the "Export" button. 3. Locate and open the generated file. | A valid, readable .csv file containing the correct alert data shall be created in the phone's local storage. |
| 2.3.4 User Profile & Recommendations Screen |  |  |  |
| REQ-019 | The app shall include a User Profile & Recommendations screen. | 1. Navigate through the app. | A screen titled "User Profile" or "Profile" shall be accessible. |
| REQ-020 | The Recommendations section shall provide actionable suggestions based on the user's current physiological state. | 1. Using a test firmware, send data indicating a "Dehydrated" state. 2. Navigate to the Profile/Recommendations screen. | A recommendation text such as "Hydration is low. Drink water." shall be displayed. |
| REQ-021 | The User Profile screen shall provide a function to un-pair the wearable device or log out. | 1. Navigate to the Profile screen. 2. Tap the "Logout" or "Un-pair" button. | The app shall either log the user out or terminate the BLE connection and remove the device from the paired list. |
| 2.3.5 Settings Screen |  |  |  |
| REQ-022 | The Settings screen shall allow the user to configure the alert thresholds. | 1. Navigate to Settings. 2. Change the value for the "High Temperature" alert threshold. 3. Save the setting. | The new value shall be saved and used by the alert engine for subsequent checks. |
| REQ-023 | The Settings screen shall allow the user to select the data synchronization interval. | 1. Navigate to Settings. 2. Change the sync interval from 5s to 10s. 3. Return to the dashboard and observe the data refresh rate. | The live data on the dashboard shall now update approximately every 10 seconds. |
| REQ-024 | The Settings screen shall allow the user to select a "User Type" profile. | 1. Navigate to Settings. 2. Change the "User Type" from "Student" to "Athlete". | The selection shall be saved and available for the algorithm to use. |